## Slide 1: Naming chemicals from Raman spectra
A model trained on public data, tested on our instrument | October 2026
Chart: 
Notes: This deck summarises the full piece of work: collecting public Raman data, analysing it, training a model to name chemicals, and testing that model on spectra from our own instrument.

## Slide 2: The result in four numbers
360
chemicals the model can name
93%
correct on public spectra it had not seen
82%
correct on our instrument (58 of 71 spectra)
6 of 26
of our chemicals exist in public data
No public Raman data exists for explosives or narcotics. To cover the chemicals we actually measure, our own spectra have to become the reference library.
Notes: The model works well on public data and transfers to our instrument for the chemicals it knows. The gap is coverage: only six of our twenty-six chemicals have any public reference spectrum.

## Slide 3: Public data covers 360 chemicals
Chemicals per public source used for training
Chart: 
| Biomolecules | Inorganic crystals | Pharma excipients | Solvents | Drug adulterants
Chemicals | 141 | 127 | 56 | 31 | 11
4,120
public spectra used, resampled to one common range
753
spectra in our own dataset, across 26 chemicals
0
public spectra for explosives or narcotics
Notes: About 1.4 GB of public Raman data was downloaded. Five sources with real chemical names were used. The RRUFF mineral database was downloaded but left out of training on request. Three chemicals appear in two sources, so the bars add up to slightly more than 360.

## Slide 4: Most chemicals have one or two spectra
Number of chemicals, by real spectra available
Chart: Real spectra per chemical
| 1 | 2 | 3 to 5 | 6 to 50 | Over 50
Chemicals | 170 | 84 | 69 | 6 | 31
1
Few examples to learn from
254 of 360 chemicals have at most two spectra, so extra training copies were generated from each one.
2
Sources are not comparable as-is
Intensity scales differ by a factor of 100,000 and only 450 to 1,800 cm-1 is covered by all.
3
Some chemicals look identical
33 pairs have nearly the same spectrum; 19 duplicate names were merged.
Notes: Only the 31 solvents have more than 100 spectra each. For the 170 chemicals with a single spectrum, accuracy on real data could not be measured, only on generated copies.

## Slide 5: Background removal is the key step
1 Resample to 450-1,800 cm-1
2 Smooth
3 Remove background
4 Square-root scale
5 Normalise
Correct first guess when the test instrument differs from the reference
Chart: 
| No background removal | Background removed | Removed + square root
FT-Raman reference, 785 nm test | 58.7 | 82.7 | 89.3
785 nm reference, FT-Raman test | 41.3 | 78.7 | 80
700 of 1,351
wavenumbers kept as the important features. Accuracy was unchanged against using all of them, and fell when fewer than 400 were kept.
Notes: Every spectrum goes through the same five steps in training and in testing. Background removal roughly doubles accuracy on a different instrument in one direction. The 700 kept wavenumbers were chosen by ranking how well each separates the chemicals.

## Slide 6: 93% correct on unseen public spectra
Correct first guess, final model
Chart: 
| Held-back public spectra (862) | Other instrument: FT-Raman to 785 nm (75) | Other instrument: 785 nm to FT-Raman (75)
Top-1 accuracy | 93.4 | 88 | 85.3
Model
Library matching combined with logistic regression
Training set
28,800 spectra: real ones plus generated copies, 80 per chemical
Model choice
Seven models compared; all within about one point, inside the noise of the test
Notes: The cross-instrument tests are the most relevant to us: the reference spectrum comes from one instrument and the test spectrum from another. The correct name was in the top five for 92 to 97 percent of spectra across the three tests.

## Slide 7: On our instrument: 58 of 71 correct
Correct first guess per chemical (number of our spectra)
Chart: 
| Sodium Nitrate (20) | Acetone (14) | Potassium Nitrate (33) | Ethyl Alcohol (2) | Potassium Carbonate (1) | Isopropyl Alcohol (1)
Correct | 100 | 100 | 73 | 0 | 0 | 0
82%
correct first guess; the right name was in the top five for 96%
Nitrates get mixed up
8 of the 9 potassium nitrate errors were another nitrate salt. All nitrates share one dominant peak near 1,050 cm-1.
Notes: Only six of our chemicals exist in the training data, giving 71 test spectra from spec_log.csv. The model had never seen a spectrum from our instrument. Leaving out the three failing chemicals gives 58 of 67, or 87 percent, but that rise comes only from removing spectra.

## Slide 8: What limits the model today
20 of 26
of our chemicals are missing
Urea, naphthalene, PNBA, hexamine and every explosive have no public reference, so the model cannot name them.
682
spectra get a confident wrong name
The model always answers. Urea is usually called iron sulfate, and RDX boric acid. It cannot yet say "unknown".
3
spectra of alcohols, all wrong
Ethyl and isopropyl alcohol failed. With so few spectra the cause is not known.
1 of 3
of our spectrum files works
Accuracy is 82% on spec_log.csv but 6% and 0% on spec_wl.csv and spec_wol.csv. Not yet investigated.
Notes: These are the reasons the model is not yet usable for our chemicals. The first is by far the largest. The model also uses only 450 to 1,800 cm-1, so information above that range is ignored.

## Slide 9: Recommended next steps
1
Add our own spectra as references
Covers urea, naphthalene, PNBA, hexamine and the explosives; test on spectra held back from each.
2
Teach the model to say "unknown"
Use our 191 "not in lib" spectra to set the decision.
3
Clarify our data files
What spec_wl.csv and spec_wol.csv contain, and what the ON / OFF column in chem_log.csv means.
4
Measure more of the rare chemicals
Seven of our chemicals have only one or two spectra.
Notes: The first step is the only route to covering the chemicals we measure. For narcotics, a commercial library would have to be licensed or reference standards measured, because no public data exists.